## Slide 1: The Airborne Cloud-Aerosol Transport System (ACATS): A status and update on development efforts at GSFC
Wind Lidar Meeting
Feb 7th
Boulder, CO
Scott Ozog
PhD Candidate
NASA Goddard Space Flight Center
University of Maryland

## Slide 2: ACATS HSRL & Doppler Wind Lidar
Fabry-Perot fringe imaging lidar at 532nm
  Rooted in heritage work from Matt McGill & Paul Hays (Univ. of Michigan)
Small enough to fit in ER-2 wing super-pod
Array of 24 detectors to resolve the spectrum and returned Doppler shift
  1.2 Free Spectral Ranges imaged
Rotating HOE used to retrieve 4 separate LOS wind vectors
  0°, 90°, 180°, and 270° at 45° off nadir angle
Optimized for retrievals in aerosol layers only
Notes: -Tech demo optimized to be built with parts at the time for the small space of the superpod
- Single wavelength because of what was available
-Emphasis on coming from Matt McGill work

## Slide 3: ACATS Lidar Investments
Leveraging investments was key to an inexpensive development.
data system
[SBIR, ESTO]
laser
[SBIR, ESTO]
etalon &
controller
[SBIR, ESTO]
HOE telescope
[SBIR, ESTO, IRAD]
receiver development
[IRAD, STTR]
circle-to-point optic
[IRAD, SBIR]
detector subsystem
[ESTO]
Notes: Slow process built from parts we had (Frankenstein): Originated summer 2008
Bits and pieces over time
Still worked well given the manner in which was built / different from dedicated $5 million to instrument and labor efforts

## Slide 4: Initial Results
ACATS first science flights were September 2012
  13 ER-2 flights out of Wallops Island, VA
Telescope rotation to all 4 looks angles worked well
  Dwell time set to 60 seconds for each look angle
  Telescope alignment worked well for first flights
    Has since been improved for each look angle
4
0o (fore)
270o (left)
90o (right)
180o (aft)
Rotates CCW
Notes: Emphasize that it worked will given its non traditional path of development

## Slide 5: ACATS HSRL & DWL
No dedicated funding
  On going side work at Goddard
  All flights have been piggybacked with CPL
    2012, 2014, 2015
Airborne tech demo test bed
  Fit in Superpod
  Seeded laser
  Fringe imaging multi-channel HSRL
  Rotating HOE
  Focus was on HSRL retrievals
    Yorks et al. 2014:
    The Airborne Cloud–Aerosol Transport System: Overview and Description of the Instrument and Retrieval Algorithms
    Worked as an HSRL
    Issues discovered with calibration and reference spectrum
      Horizontal wind algorithms never fully tested
Flies in ER-2 Superpod

## Slide 6: ACATS Updates/Improvements
Past few years have had dedicated PhD student working on ACATS
New hardware integrated into ACATS telescope (IRAD 2017)
  Improved consistency and reliability of etalon calibration
  Improved retrieval of laser reference spectra
ACATS has been operating at NASA GSFC collecting zenith pointing data of clouds and aerosols for algorithm and hardware testing

## Slide 7: Improved Etalon Calibration
Old Calibration Method
Old Calibration
New Calibration
Poor Fit
Accurate Fit
Notes: After demo flights, worked, but now well. Sat around with no dedicated funds and campaigns

## Slide 8: Improved Etalon Calibration
Old Calibration Method
Old Calibration Consistency
New Calibration Consistency

## Slide 9: ACATS Wind Test Data
Notes: 60 second integration for reference spectrum
360 second (6min) interval
Focus on and reiterate this is up looking, not flight data, looking at vert wind, expected zero wind

## Slide 10: ACATS Wind Test Data
Laser reference spectra were retrieved at 6 minute intervals
Spectra created from 60 second integration period
  Future testing will be done to optimize these specs
Notes: This slide my be taken out. Information can be given on previous slide for time sake

## Slide 11: ACATS Vertical Wind Profiles
Vertical Resolution: 0.6 km Horizontal Resolution: 25 seconds (5 km airborne)
Notes: Emphasis on limited to pointing zenith, wind should be zero, here’s our test
Looking into layer only, ignoring overlap region from the ground (not problem plane) emphasis again ACATS is optimized for aerosol layer only
19:34:33 Mean = -0.37 m/s
20:30:07 Mean = -0.58 m/s

## Slide 12: Future Work & Conclusion
No dedicated funding source for ACATS
  Is wind lidar work going on at GSFC
Still undergoing algorithm updates and testing
  Improve wind retrievals
  Improve HSRL products
Next planned flights Summer 2018 out of Palmdale, CA
  Piggyback (like all previous flights)
  Test updated etalon calibration and horizontal wind retrievals from airborne environment
Further plans
  Science tool in future field campaigns
  Potential cal/val instrument
  Stepping stone to future space missions
Notes: What if we had money?
-Could buy flight hours, dedicated personal, build new receiver and/or etalon for Rayleigh, etc. Could do lots with dedicated funding

## Slide 13: Back Up

## Slide 14: Back Up
A Fabry-Perot interferometer is used to
resolve the spectral signature.
Wavelength is a function of radius
in the image plane.
The challenge is to efficiently measure
the annular image.
normalized intensity
radial displacement
14

## Slide 15: Back Up
Use of a Fabry-Perot interferometer results in different system response for the aerosol and molecular components of the total signal. In fact, the molecular return is imaged as a nearly flat spectrum that can be subtracted from the total signal.

## Slide 16: Back Up
Our solution: use a custom holographic optical element with a linear array detector to measure the fringe pattern.
U.S. Patent #631908

## Slide 17: Back Up
Mounted servos
Delrin
Filter over
fiber head

## Slide 18: Back Up
laser type | Nd:YAG, seeded
operating wavelength | 532 nm
laser repetition rate | 200 Hz
laser output energy | £ 10 mJ/pulse
laser beam divergence | 100 mrad, full angle
telescope diameter | 8 inches
viewing angle | 45 degrees
telescope field of view | 300 mrad, full angle
bandpass filter | 150 pm FWHH
etalon spacing | 10 cm
etalon reflectivity | 85%
orders imaged | 1.2
detector channels | 24
measurement dynamic range | 480 m/s (L-O-S)
platform altitude | ~ 20 km (65,000 ft)
platform speed | ~ 200 m/s
wind error | 2 m/s (L-O-S)
aerosol extinction error | 0.025 km-1
vertical resolution (wind) | 0.5 – 1.0 km
vertical resolution (aerosol ext.) | 300-600 m
horizontal resolution (wind) | 25 seconds (~ 5 km)
horizontal resolution (aerosol ext) | 25 seconds (~ 5 km)

## Slide 19: Etalon Scan
During calibration ACATS software changes the etalon plate spacing in short steps (pm), which moves the fringe pattern (peak) across all 24 channels (etalon scan)
Each channel’s response during scan is measured
Measured Response Channel 1
Scan Step
Photon Counts
ACATS Spectrum During Scan
Notes: Calibration also known as an etalon scan

## Slide 20: Etalon Transmission Fit
Etalon transmission function and transmission parameter
Iterative least squares fitting technique is applied in post processing to determine what calibration value (defect value) yields the best fit
Notes: -Here we have the transmission function for an etalon which is a function of its plate spacing, # channels, wavelength, and defect
-Defect starts at 0 and increased by intervals of 1 to determine what value yields best fit

## Slide 21: ACATS Lidar Equation

## Slide 22: ACATS Least Squares Retrieval
Re-write the equation in matrix form:
Same equation can be written more generally as:
Use weighted least-squares fit to find solutions for 3 parameters
  W is weighting matrix
  G is generalized matrix to be inverted

## Slide 23: ACATS Algorithms: Wind
Wind products
  Laser tilted 45o to measure LOS component of wind
  Use scanning scheme, assuming that the wind field over the sensed volume is horizontally homogeneous
  Flight software directs ACATS to stop and stare at four look angle positions (azimuth aircraft-relative angles):
  2 orthogonal LOS measurements are corrected for aircraft motion and used to compute horizontal wind (Lee et al. 1994; Leon and Vali 1998)
0o (fore)
270o (left)
90o (right)
180o (aft)
Rotates Counter-Clockwise

## Slide 25: Initial HSRL Results
Standard vs HSRL:
Agreement within 10% for ATB computed using HSRL method (red) and ATB using standard lidar technique (blue)
17