## Slide 1: TWiLiTE Operations During Polar Winds and Analysis Status
Bruce Gentry1, Huailin Chen2 and Steve Mitchell3
1NASA/GSFC
2SSAI
3SigmaSpace Corp
Space Wind Lidar Working Group
Miami, FL
November 4, 2015

## Slide 2
TWiLiTE is a compact, molecular direct detection scanning Doppler lidar designed to measure wind profiles in clear air from 20 km to the surface.
TWiLiTE is modular and operates autonomously on NASA research aircraft (ER-2, DC-8, WB-57, Global Hawk).
Engineering flight tests on the NASA ER-2 in 2009, 2011 and 2012 demonstrated autonomous operation of all major systems.
TWiLiTE is now installed in the Nadir 2 port of the DC-8.
Tropospheric Wind Lidar Technology Experiment – TWiLiTE
TWiLiTE system mounted in the DC-8
TWiLiTE Performance Summary
Notes: TWiLiTE was developed as an IIP program funded by NASA ESTO. Additional support from the AITT program allowed integration and flight testing on the NASA ER-2.
TWiLiTE Supports the 3DWinds NRC Decadal Survey Mission. It represents the direct detection molecular or ‘clear air’ portion of the recommended ‘Hybrid Doppler lidar concept’. The coherent Doppler lidar utilizes aerosol backscatter a complementery measurement capability and is the other half of the Hybrid concept.
(Lower left panel) The TWiLiTE Doppler lidar operates at 355 nm in the uv and measures wind profiles by measuring the Doppler shift of the laser backscattered return from molecules in the atmosphere using a ‘double edge’ etalon receiver (similar to the ADM-Aeolus molecular channel).
Designed to measure full profiles of horizontal winds from the flight altitude of the NASA ER2 or Global Hawk (20 km) to the surface with 250 m vertical resolution and <2 m/s accuracy
Advances TRL of key technologies identified for future space missions e.g the single frequency uv NdYAG laser; the molecular Doppler receiver; Tunable Fabry-Perot etalon filter and the rotating conically scanned telescope.
These technologies have all been demonstrated at the system level in engineering flights of TWiLiTE on the NASA ER-2 high altitude aircraft in 2009, 2011 and 2012.
TWiLiTE was reconfigured, but not flown, for the NASA Global Hawk as part of the Hurricane and Severe Storm Sentinel (HS3) Venture Class Mission.
Small modifications to the instrument are underway to allow integration on the NASA DC-8 in the Nadir Port 2 location (DAWN goes in Nadir 7).
DC-8 flights are planned in February, 2015 to validate TWiLiTE performance on the DC-8

## Slide 3
TWiLiTE Compatible
NASA Airborne Science Platforms
DC8
WB57
ER2
25
20
15
10
5
Max Altitude (km)
6-8 hrs duration
Unattended operation
Global Hawk
36 hrs duration
Unmanned vehicle
TWiLiTE configured for DC8 Nadir Port 2
(AITT)
TWiLiTE configured for WB57 3’ Pallet
(ESTO IIP04)
TWiLiTE configured for ER-2 Q-Bay
(ESTO IIP04)
TWiLiTE configured for Global Hawk Zone 25
(HS3)

## Slide 4
DC-8 Operator Console
Single rack with CPU, monitor and KB. Connect to TWiLiTE directly via Ethernet cable.
Command screen with log on the right and housekeeping data on the left

## Slide 5
Console – Etalon calibration scan

## Slide 6
Console – Real time 7-channel PMT signal display

## Slide 7: TWiLiTE Operation Summary Polar Winds- May 9-18, 2015
Date | Scan Cycles | Etalon Calib Sweeps | 
May 9 | 77 | 14 | Transit to Keflavik
May 11 | 44 | 10 | Etalon T=3 deg C
May 13 | 39 | 8 | Etalon T=7 deg C
May 15 | 46 | 9 | Etalon T=9 deg C
May 16 | 71 | 17 | Etalon T=13 deg C
May 17 | 101 | 10 | Etalon T=6 deg C
May 19 | 62 | 7 | Etalon T=3 deg C
May 21 | 52 | 10 | Etalon T=3 deg C
May 23 | -- | -- | HD error
May 24 | -- | -- | HD error
May 25 | 36 | 7 | Etalon T=7 deg C
May 28 | 109 | 14 | Transit back to DFRC

## Slide 8: Step Stare Azimuth Scanning and Sampling
45
90
135
225
270
315
6 direction step stare scanning.
The HOE is rotated CCW to each of the 6 directions. Dwell at each direction is 30 sec. ~3 sec to move and settle
Aircraft heading
Shot pattern in Earth-relative rotated
reference frame.

## Slide 9
May 15, 2015 Polar Winds Flight track with TWiLiTE scans
Complete scans: 46
Partial scans: 2
Etalon cals: 9

## Slide 10: Edge 1 – Hi Signal ~ 90% May 15, 2015 Cycle 26

## Slide 11: Edge 1 – Med Signal ~ 9% May 15, 2015 Cycle 26

## Slide 12: Edge 1 – Lo Signal ~1% May 15, 2015 Cycle 26

## Slide 13: EM Channel – No etalon filters

## Slide 14: Edge 1 – Hi, Edge Ratio May 15, 2015 Cycle 11
Zavg=100 m
Tavg=30 s

## Slide 15: Edge 1 – Med, Edge Ratio May 15, 2015 Cycle 11
Zavg=100 m
Tavg=30 s

## Slide 16: Edge 1 – Lo, Edge Ratio May 15, 2015 Cycle 11
Zavg=100 m
Tavg=30 s

## Slide 17: May 27, 2015: TWiLiTE Processing Status
All 7 PMT channel signals have been processed for background correction, dead time correction.
Signal files have been time synched and merged with DC-8 navigation data and corrected for pitch, roll and instrument offsets determined from ocean surface returns at the six angles.
All etalon scan data have been analyzed and fitted to the instrument function.
Currently examining the Edge channel ratio data (proportional to LOS velocity: wind+aircraft). There are some issues with the etalon alignment drift due to temperature. There are also indications that the aircraft velocity pushes the dynamic range of the Edge filters particularly for the aft looking directions (135 deg and 225 deg azimuth). Additional analysis is required particularly for the etalon alignment drift.
Next steps-
  develop more detailed model of temperature effects on the etalon alignment and incorporate this into the calibration algorithms.
  process radial LOS wind data for the 6 directions.
  Generate u,v component wind profiles

## Slide 18
May 9, 2015 Polar Winds Transit Flight track with TWiLiTE scans
Palmdale, CA to Keflavik, Iceland
Complete scans: 77
Partial scans: 3
Etalon cals: 14

## Slide 19: Internal Etalon Temperatures May 9, 2015 Transit Flight

## Slide 20: Etalon Gap Drift May 9, 2015

## Slide 21
May 15, 2015 Polar Winds Flight track with TWiLiTE scans
Complete scans: 46
Partial scans: 2
Etalon cals: 9

## Slide 22: Internal Etalon Temperatures May 15, 2015

## Slide 23: Etalon Gap Drift May 17, 2015

## Slide 24: Internal Etalon Temperatures PECAN June 19, 2015

## Slide 25: New Calibration Procedure – 2
Start Etalon Scan
End Etalon Scan
Data Collection
1. Fit Start and Ending Etalon Scans for each data collection period

## Slide 26: PECAN June 19, 2015

## Slide 27: New Calibration Procedure – 1
Start Etalon Scan
End Etalon Scan
Data Collection
1. Convlove Start and Ending Etalon Scans with RBS lineshape for range of T,P

## Slide 28: New Calibration Procedure – 3
Start Etalon Scan
End Etalon Scan
Data Collection
1. Generate calibration curves 2*(E1-E2)/(E1+E2) for each etalon scan

## Slide 29: New Calibration Procedure – 4
Interpolate between Start and Ending Calibration curves for each 1 s profile
within the data collection period

## Slide 30: Summary
Thermal environment presented major challenge for TWiLiTE. Etalon drifts of order 200 to 300 m/s equivalent velocity were observed as etalon slowly reached thermal equilibrium during flights.
Assumptions of etalon stability between periodic (~every 15 minutes) calibration scans, previously used in the TWiLiTE quick look analysis, are invalid for Polar Winds flights.
A new calibration procedure to use all etalon scans to dynamically calibrate each 1 sec profile.
This procedure has been successfully used in the analysis of TWiLiTE data collected from the ground during the PECAN field campaign in June-July, 2015
Currently implementing this procedure in the TWiLiTE airborne data analysis.

## Slide 31: Next Steps
Continue implementation of the dynamic calibration procedure to correct for thermal etalon drift and mis-alignment
Produce corrected first look winds for all flights
Investigate use of Energy Monitor channel for aerosol and cloud measurements
Investigate use of Edge channel signals for determination of the atmospheric temperature.
Consider lessons learned from Polar Winds and PECAN to improve operations, especially thermal control, for future missions.