## Slide 1: ALADIN Airborne Demonstrator (A2D) – first results from WindVal campaign 2015
Uwe Marksteiner, Oliver Reitebuch, Christian Lemmerz, Benjamin Witschas, Stephan Rahm, Andreas Schäfler, Engelbert Nagel

## Slide 2
Recall of 2009 results
First results from 2015 airborne campaign in Iceland
► Outline
  Calibration
  Winds

## Slide 3
The A2D is a prototype of the satellite instrument and is operated by DLR from ground and as an airborne demonstrator. Two different types of interferometers analyse Rayleigh and Mie scattering from molecules and aerosols or clouds.
► The A2D subsystems

## Slide 4
► Airborne campaign 2009
calibrations
wind observations
1st time worldwide 2 wind lidar instruments measuring simultaneously on same aircraft
A2D (direct-detection) and 2-µm (heterodyne)
2-µm: low random & systematic errors  reference lidar
Notes: For the first time worldwide two wind lidar instruments were operated simultaneously on the same aircraft in 2009: The A2D and a well-established coherent reference lidar at 2 µm, the latter being characterized by low random and systematic errors of about 0.1 m/s. During the campaign in 2015 two additional lidar systems also using both detection principles were operated in paralled on a NASA DC-8 aircraft.

## Slide 5
LOS wind observations on 2009/10/01
2-µm
Mie
Rayleigh
ECMWF
MODIS, Oct. 1st , 2009
► Airborne observations and ECMWF model

## Slide 6
Mie
Rayleigh
slope: 1.14
std. dev.: 1.5 m/s
# of points: 932
slope: 1.04
std. dev.: 2.5 m/s
# of points: 596
► Statistical Comparison for ALADIN airborne demonstrator

## Slide 7
vZWC: Zero Wind Correction value
vac: aircraft velocity
vGR: ground return velocity
vmeas: correction of measured wind speed
 vwind: true wind speed
2009-09-29
blue: wind meas.
red: for ZWC
objective to validate method of ZWC achieved
correction by mean value needed instead of observational basis
ZWC values different for Mie and Rayleigh
► Zero Wind Correction
blue: measured
red: corrected
2-µm
A2D
Rayleigh

## Slide 8
Recall of 2009 results
  Winds
► Outline
First results from 2015 airborne campaign in Iceland
  Calibration

## Slide 9
► scientific flights & objectives
ASCAT underflight
Iceland & southerly jet
2 calibrations
2 calibrations
over sea ice
1 calibration
Tip-Jet
jet stream
TDS underflight
Greenland
summit station

## Slide 10
► calibration locations 2009 & 2015
2 calibrations
over sea ice
1 calibration
2 calibrations
2009
2015
2 calibrations

## Slide 11
► A2D calibration - Rayleigh channel (movie)

## Slide 12
I (B)
I (A)
Wind
Frequency
Response:
► A2D calibration - Rayleigh channel

## Slide 13
4 × auxiliary range gates (L0-L3)
1 × internal reference (INT=L4)
► Rayleigh response curves (2015-05-16 a)
4 × auxiliary range gates (L0-L3)
1 × internal reference (INT=L4)
20 × atmospheric range gates (L5-L24)
4 × auxiliary range gates (L0-L3)
1 × internal reference (INT=L4)
20 × atmospheric range gates (L5-L24)
2 × ground return (L19-L20)
Notes: RAY responses per layer for selected pix. / LOSV corr.ected / TLE frequencies

## Slide 14
► Mie response curves (2009-09-21 b)
only Internal Reference & ground return

## Slide 15
TLE: commanded relative frequencies
⇒ Aeolus & A2D
WM: measured absolute frequencies
⇒ only A2D
► Rayleigh calibrations – internal reference (WM corrected)
Cal. | slope / MHz-1 | intercept / -
1 | 0.000450 | -0.00070
2 | 0.000450 | -0.00096
3 | 0.000458 | -0.00090
4 | 0.000457 | -0.00019
5 | 0.000455 | -0.00133
6 | more QC needed | 
7 | 0.000456 | -0.00083

## Slide 16
► Rayleigh calibrations – ground return (WM corrected)
Cal. | slope / MHz-1 | intercept / -
1 | 0.000458 | 0.01582
2 | 0.000460 | 0.01592
3 | 0.000455 | 0.01300
4 | 0.000454 | 0.01444
5 | 0.000452 | 0.01025
6 | more QC needed | 
7 | 0.000457 | 0.00947

## Slide 17: ► Mie calibrations – internal reference (TLE)
Cal. | slope / MHz*pix-1 | intercept / pix
1 | -98.85 | 7.30
2 | -98.37 | 7.29
3 | -99.79 | 7.40
4 | -99.94 | 7.41
5 | -98.78 | 7.42
6 | more QC needed | 
7 | -98.49 | 7.37
Notes: „wrt Mie centre“  center frequency is the one of the step with the response closest to 7.5 (same intensity on pixel 7 & 8)

## Slide 18: ► Mie calibrations – ground return (TLE, sum, LOSV corr.)
Cal. | slope / MHz*pix-1 | intercept / pix
1 | -97.21 | 7.21
2 | -96.75 | 7.24
3 | -96.61 | 7.29
4 | -97.08 | 7.31
5 | -95.99 | 7.32
6 | more QC needed | 
7 | -94.90 | 7.19

## Slide 19
Recall of 2009 results
  Calibration
► Outline
First results from 2015 airborne campaign in Iceland
  Winds

## Slide 20
distributed
ground echo
ground speed ≈ 0 m/s
(sea surface)
Internal Reference (0 m/s)
flight altitude 9100 m
► 2015-05-19 - Greenland: A2D Rayleigh LOS winds
clouds

## Slide 21
cloud
Mie winds match well
with Rayleigh wind field
ground speed ≈ 0 m/s
(sea surface)
altitude ≈ 1500 m
altitude ≈ 6300 m
► 2015-05-19: Greenland: valid A2D Mie LOS winds
Notes: Rayleigh winds processed with case 33: Greenland 20150516s2 - Cal.3 (TLE/+-825/DCO checked/ERR 5)

## Slide 22
A2D
2-µm
► 2015-05-19: Rayleigh winds A2D  2-µm
Notes: Rayleigh winds processed with case 39: Greenland 20150516s2 - Cal.3 (WM/sym.1500/DCO checked/ERR 5)

## Slide 23
► 2015-05-19: Rayleigh winds A2D  2-µm (Cal.7)
Layers 8 – 20 (INT = 4)
coverage threshold = 85%
Statistical results Cal.3 Cal.7
corr. coeff. r 0.90 0.87
N points 759 758
slope 1.27 1.15
intercept (m/s) -2.53 -1.30
std. dev. (m/s) 2.15 2.07
avg. bias (m/s) 0.22 -0.03

## Slide 24
► Conclusion
more than 12 hours of wind measurement data from 2015 campaign
similar performance in wind measurement as in 2009
calibrations make the difference
temperature/pressure correction will be investigated using DC-8 dropsondes
more quality control steps still to be included in wind analysis
a lot of manual adaptation and optimization necessary

## Slide 25
BACK UP

## Slide 26: ► Rayleigh calibration profiles (altitude wrt instrument)
Cal.3 & 4 at
same spot
Cal.1 & 2 at
same spot
Cal.5 deviates
despite same
location as
Cal.3 & 4!
internal reference
ground
return
atmosphere
Cal.1: 2009/09/21 over Greenland ice sheet
Cal.2: 2009/09/21 over Greenland ice sheet
Cal.3: 2015/05/16 a over Greenland ice sheet
Cal.4: 2015/05/16 b over Greenland ice sheet
Cal.5: 2015/05/16 c over Greenland ice sheet
Cal.6: 2015/05/23 a over sea ice (needs more QC)
Cal.7: 2015/05/23 b over sea ice
extrapolated
values
below ground
sea ice
20 MHz
2%