## Slide 1: Results from the ATHENA-OAWL Venture Tech Flights
Sunil Baidar1,2, S. Tucker3, M. Beaubien4, R. M. Hardesty1,2
1. CIRES, University of Colorado Boulder, Boulder, CO
2. NOAA Earth System Research Laboratory, Boulder, CO
3. Ball Aerospace & Technologies Corp., Boulder, CO
4. Yankee Environmental Systems, Turner Falls, MA

## Slide 2: Questions
How well does the GrOAWL instrument measures wind?
  What is the accuracy?
  What is the precision?
What was the instrument performance during the flights?
What can we say about the ATHENA-OAWL performance based on the aircraft measurements?

## Slide 3: GrOAWL Instrument
Airborne demonstrator for ATHENA-OAWL
Aerosol based direct detection Doppler wind lidar
Operates at 532 nm
Uses a field-widened, quadrature channel, Mach Zehnder interferometer
Two azimuthally orthogonal Line of Sights
  FWD look: 45° Elevation angle and 45° Azimuth angle
  AFT look: 45° Elevation angle and 135° Azimuth angle
Remotely operated
Laser Electronics
Fiber coupled OAWL Interferometer
Payload Controller
Laser-2
GPS/ IMU
  Variable 532 nm output power
Laser-1
  2-looks, 2 telescopes

## Slide 4: GrOAWL AOVT flights
Deployed on NASA WB-57
May-June 2016
8 flights
  5 engineering
  3 validation
Racetrack patterns over the Gulf of Mexico
  Revisit times: ~1hr/loop
  Launched dropsondes from the aircraft

## Slide 5: Validation Efforts
Validation focused on sondes
  High Definition Sounding System (HDSS) dropsondes
    Yankee Environmental Systems with support from Office of Naval Research
    Automated dropsonde system on WB-57
    Dropped 40+ dropsondes
  Balloon-sondes launched from NOAA NWS station at Corpus Christi
NOAA High Resolution Rapid Refresh (HRRR) model winds
  3 km resolution, hourly updated
  Radar data assimilated every 15 min over a 1-hour period
  Comparison with 0 hour forecast/analysis winds

## Slide 6: GrOAWL data
Look 1 (FWD)
Look 2 (AFT)
Time resolution: 10 seconds
Vertical resolution: ~120 m

## Slide 7: Dropsonde Comparison
No spatial overlap with dropsondes
Closest GrOAWL measurement in space to dropsondes
10 seconds profiles within 1 minute
Project dropsonde winds to GrOAWL LOS
Notes: Change profile color coded by spatial difference.

## Slide 8: Wind Profiles: June 17 2016
Combine two looks to retrieve wind speed and direction
Assume vertical velocity = 0.
8 km
16 km, ~2 mins
Notes: Add a slide with only the GrOAWL data. (like animation? 60

## Slide 9: Wind Profiles: June 21 2016
Combine two looks to retrieve wind speed and direction
Assume vertical velocity = 0.
8 km
16 km, ~2 mins
Notes: 5

## Slide 10: Spatial Variability
GrOAWL
HRRR
GrOAWL measurements captured the spatial variability predicted by HRRR.

## Slide 11: What is the accuracy of the GrOAWL wind measurements?

## Slide 12: Accuracy
21 total dropsondes from 3 validation flights
Dropsonde data is considered the “True Value”
10 seconds data

## Slide 13: Accuracy
21 total dropsondes from 3 validation flights
Dropsonde data is considered the “True Value”
Many 10 seconds data to better estimate the GrOAWL accuracy (1 minute)

## Slide 14: Accuracy
21 total dropsondes from 3 validation flights
Dropsonde data is considered the “True Value”
Many 10 seconds data to better estimate the GrOAWL accuracy (5 minutes)

## Slide 15: Wind Speed and Direction
Wind Speed
Wind Direction
Very good agreement between GrOAWL and Dropsondes.
Notes: 30

## Slide 16: What is the precision of the GrOAWL wind measurements?

## Slide 17: Precision
Depends on
  Aerosol conditions
  Range
  Accumulation/Averaging
Precision for 10 seconds profiles
  CNR based

## Slide 18: Precision: Time series

## Slide 19: Precision: Time series

## Slide 20: Precision: Time series

## Slide 21: Precision
Depends on
  Aerosol conditions
  Range
  Accumulation/Averaging
Precision for 10 seconds profiles
  CNR based
CNR based precision estimate agrees very well with precision from time series analysis

## Slide 22: What was the instrument performance during the flights?

## Slide 23: GrOAWL Performance
Detected Photons
Molecular + Aerosol photons
Temperature and pressure measured by the dropsonde
Notes: Ground based measurement along with NCAR HSRL instrument is ongoing. – to better understand the instrument performance under low aerosol conditions.

## Slide 24: GrOAWL Performance
Detected Photons
Overlap function is complicated
Focus on the lowest 5 km
Notes: Ground based measurement along with NCAR HSRL instrument is ongoing. – to better understand the instrument performance under low aerosol conditions.

## Slide 25: GrOAWL Performance
Detected Photons
Ground based measurement along with NCAR HSRL -- ongoing
~45% reduction in instrument throughput.
Cause:
Aircraft vibration
Window fogging
Notes: Ground based measurement along with NCAR HSRL instrument is ongoing. – to better understand the instrument performance under low aerosol conditions.

## Slide 26: GrOAWL Performance
Instrument Contrast
Good understanding of the instrument performance during AOVT flights.
Ability to resolve fringes
~55% reduction in instrument contrast
Cause:
Aircraft vibration.
Notes: Ground based measurement along with NCAR HSRL instrument is ongoing. – to better understand the instrument performance under low aerosol conditions.

## Slide 27: GrOAWL Performance
Detected Photons
Velocity Precision
Instrument performance by lower by factor of ~4.
Notes: Ground based measurement along with NCAR HSRL instrument is ongoing. – to better understand the instrument performance under low aerosol conditions.

## Slide 28: What can we say about the ATHENA-OAWL performance?

## Slide 29: GrOAWL vs ATHENA-OAWL
Main Differences:
Energy
1.5 mJ vs 45 mJ
Telescope size
0.29 m vs 0.7 m
Range
10 km vs 400 km
Accumulation Volume Height
~120 m vs 1000 m
ATHENA-OAWL performance is expected to be similar to the GrOAWL performance during AOVT flights in the lower atmosphere.

## Slide 30: GrOAWL vs ATHENA-OAWL
ATHENA-OAWL performance is expected to be similar to the GrOAWL performance during AOVT flights in the lower atmosphere.

## Slide 31: Summary
Comparison between GrOAWL measured winds and dropsonde winds show excellent agreement.
  Accuracy
GrOAWL precision is very good (but depends upon aerosol conditions).
Sensitivity was down (4x) but the instrument performance is well understood.
  Ground based measurement along with NCAR HSRL -- ongoing
Space-borne ATHENA-OAWL system is expected to perform very similar to the GrOAWL during the flights in the lower atmosphere.
Successfully demonstrated the ATHENA-OAWL two look concept
  Airborne instrument for future science studies
  Validated concept for space-based system

## Slide 32: Thanks to:
ESSP and ESTO for the Venture-Technology funding
The Ball GrOAWL engineering team
Fibertek (Floyd Hovis)
YES (Mark Beaubien) & ONR for dropsonde validation support
NWS Corpus Christi office for the sonde launches
The engineering team, crew, pilots, and program staff of the WB-57 for all their excellent flight support